FCW  Series
INSTRUCTION MANUAL

[image: image5.wmf]SET

UNIT

CAP

3

DIVS

0.5

AOFF

TARE

UNIT

CAL

t.b.r

GrAn

4800

9600

ON

OFF

Unit

Unit

Unit

Unit

Unit

Unit

Unit

Unit

Unit

Unit

TARE

UNIT

(FCS3.1)

TARE

UNIT

OVL

TARE

UNIT

TARE

UNIT

TARE

UNIT

TARE

UNIT

TARE

UNIT

2800

HOLD

2.5

3

HOLD

HOLD

1

2

trd

TARE

UNIT

HOLD

ON

OFF

HOLD

HOLD

TrAn

ON

OFF

TARE

UNIT

0

000

U

N

I

T

Press [TARE],flash digit move to right;

Press  [HOLD],flash digit increases.

BL

TARE

UNIT

Unit

HOLD

HOLD

120

180

240

360

OFF

HOLD

HOLD

HOLD

HOLD

HOLD

ON 

kg

OFF 

kg

ON 

lb

OFF 

lb

ON 

oz

OFF 

oz

TARE

TARE

T

A

R

E

15

30

60

120

OFF

HOLD

HOLD

HOLD

HOLD

HOLD

HOLD

HOLD

HOLD

HOLD

HOLD

Press the UNITS + TARE key

Turn the power on


KEY DESCRIPTIONS：
	ZERO
	UNITS
	TARE
	HOLD


1.  ZERO：
This key is “Zero” feature if load is below 4% of full capacity or a negative value
2.  UNITS：

A. Select weigh units：kg or lb .

3. TARE:

A. When there’s load on the scale , press this key to return to zero.

B. Put a container on the scale, press this key to tare the weight, and the tarred value is stored.

4.  HOLD：

A. Press this key to store weight value.
5.   POWER SWITCH :when the scale is used or not used ,please turn the power on or  the power off.

1
TABLE =NET/GROSS KEY)

[image: image6.wmf]
6
interface：

1：signal interface：（microphone socket）

          Pin 1（red）-------(E+（positive source）

          Pin 2（blue）-------(S+（positive signal）

          Pin 3（black）-------(E-（negative source）

          Pin 4 (white)-------(S-（negative signal）

          Pin 5（no）

          no sign        -------(GND

2：Rechargeable battery socket：

          AC/DC 9V――――――inner positive ,external negative
3：RS232 interface：DB9
RS232 TRANSMIT FUNCTION

1.  When the function trAn set ON，The RS232  is actived 
2.  Baud rate is selected 4800 or 9600 when the function t.b.r. set 4800 or 9600 。
3.  RS232 Transmission Agreement：

· Mode：Simplex Asynchronous Serial
· Baud rate：4800 or 9600
· Data Bit：8

· Parity Bit：NONE

· Stop Bit：1

· Data Format：ASCII

4.   Transmission Information Format： 16 Byte ，blank=20H

<STX><Status><+/-><D><D><D><D><D><D><D><U><U><U><U><U><ETX>

5.   Transmission Information Format show:
STX: Start of Transmission=02H

Status：Status bit

Example：*=(2AH)=Stable weight, ~=(7EH)=Unstable Weight

+/-: +=(2BH)=Positive Weight , -=(2Dh)=Negat Weight

DDDDDDD: Value of Weight , 7 ASCII Code

UUUUU: 4 ASCII Code Weight Units: kg or LB

ETX: Stop of Transmission =03H
5
THE MEANING OF FUNCTION SETTING

1． AOFF：  Time for auto off (120s / 180s / 240s /300s / off)
2． bL ：    Back Light (15 / 30 / …. / off)
3.   Unit     Weight units ( kg / lb/oz）
4． CAP：   Capacity (3/6/15/30/60/150/……/2800kg )
5． OvL：   Over load (2.5/3 )
6． CAL：   Calibration capacity (1/4，1/2，full)
7． trAn：   Transfer (On/Off)
8． t.b.r.：   Transfer band rate (4800/9600)

9.  [image: image1.png]


:    Gravity(On/Off)
    10.  [image: image2.png]


:    Transfer mode (On/Off)

SETTING THE ACCELERATION TO GRAVITY MODE

1. Press [TARE] and [UNITS] first, then turn the power on, while keeping [TARE] and [UNITS]pressed, the display will show “SEt”, then press the [UNITS] several times until the display show ”GrAv” press [TARE] the display show “on” press the [HOLD] the display show “OFF” .When the display show “ON” means acceleration to gravity is actived. When the display show “OFF” means acceleration to gravity is inactived.

2. When acceleration to gravity is actived turn the power off, then turn the power on,  then press [UNITS] key 3 seconds enter to acceleration to gravity mode .1). when the display show “C9.○ 00” ( ○ for flash digit ) first setting the local gravity.2). press [TARE] key again, rightward flash digit, 3).press [HOLD ] key, increase digit input the gravity value, press [UNITS] the display enter to another local acceleration to gravity mode.
3. When the display show  “U9.○ 00” ( ○ for flash digit) repeat the steps 1)-3) .press [UNITS] the display will return the normal weighing mode. 
2

Features：

1.   Compact Dimensions：950mm x 500mm x 50mm
2.   Material：Aluminum Housing / SUS 304 Base Cover
3.   Display options：5 digital LCD

4. Capacity/Division/ Calibration Weight  

Table 1
	CAPACITY
	DIVISION1
	DIVISION2
	DIVISION3

	2.5/3kg
	0.5g
	1g
	2g

	5/6kg
	1g
	2g
	5g

	10/15kg
	0.002kg
	0.005kg
	0.01kg

	25/30kg
	0.005kg
	0.01kg
	0.02kg

	50/60kg
	0.01kg
	0.02kg
	0.05kg

	100/150kg
	0.02kg
	0.05kg
	0.1kg

	250/300kg
	0.05kg
	0.1kg
	0.2kg

	500/600kg
	0.1kg
	0.2kg
	0.5kg

	1000/1500kg
	0.2kg
	0.5kg
	1kg

	2500/2800kg
	0.5kg
	1kg
	2kg


5.   Power：recharge / AC adapter 12V 500mA [image: image3.jpg]:Gi



 
6.   Auto back light (option) 

7. Power saving（auto shut-off timing selectable：120seconds/180seconds/240seconds/300seconds/Off）
8.   Operating Temperature：5～35℃ 

9.   Operating Humidity：25％～95％RH
10.  4Keys：  ZERO, UNITS, TARE, HOLD
ZERO        This key is “Zero” feature if load is below 4% of full capacity or a negative value
UNITS         Toggle among KG or LB unit
TARE         Tare the weight
HOLD         Press this key to store weight value
3
11.   Zero Range (4% to full capacity) 

12.   Tare Range (100% to full capacity) 

13.   Error Message indication：

              EEEEE   Ovrload

       ErrL     Power on zero count too low

       ErrH     Power on zero count too high

       ErrE     Eeprom Error

    [image: image4.png]


      Low battery indication
14. Two calibration weight units : kg , lb,oz.
15.   Open calibration capacity 
16.   Power on zero-setting range: + 10%  

17.   Zero range: + 4% of full capacity 

18.   Three modes : Normal mode / Setting mode /Internal Auto Calibration mode 

19.   RS232 function
When to calibrate
Calibration may be required when it is initially installed, if the scale/balance is moved to a substantial distance.
This is necessary because the weight of a mass in one location is not necessarily the same in another location. Also, with time and use, mechanical deviations may occur.
Calibration Weight

1. Turn the power off. Press [UNITS] and [ZERO] first, then turn the power on., while keeping pressing [UNITS] and [TARE], the display will show the “CAL  ukg”.

2. Press [ZERO] key, the display will show AD value, press [ZERO] key, the display will show “SAVE”, and then show calibration value, place the calibration value, press [ZERO] key, the display will show “SAVE”, and then return to AD value.
3. Now calibration weight is complete.
4
